## Slide 1: Blinking Orbital Prosthesis
Carmen Coddington
Bryan Jepson
Elise Larson
Michelle Tutkowski
Greg Gion-Medical Art Prosthetics
Willis Tompkins- BME Department

## Slide 2: Points of Interest
Client Information
Background
Problem Statement
Design Criteria
Competition
Design Alternatives
Final Design
Future Work
Questions

## Slide 3
Greg Gion
Medical Art Prosthetics Clinic
  Prosthetic ears, noses, eyes, fingers, hands, etc.
  Provide most life-like restoration as possible
Client Information
http://www.medicalartprosthetics.com/content.php?page=prostheses

## Slide 4: Background
Many causes for eye loss
Orbital prosthetics give a more natural appearance
Prosthetics do not blink
http://www.medicalartprosthetics.com/content.php?page=prostheses&sec=orbital

## Slide 5: Problem Statement
Orbital prosthesis that blinks on command
  Blinks reliably
  Primarily self-contained
  External controller
  Natural appearance
http://www.medicalartprosthetics.com/content.php?page=prostheses&sec=orbital

## Slide 6: Design Criteria
Blinks reliably on command
Contained in 5.5 cm3
Weigh <45 g
No detrimental health effects
Maintained easily
$500 budget

## Slide 7: Competition
http://www.sciencedirect.com.ezproxy.
http://www.freepatentsonline.com/20020049023.pdf
Electromagnetic Activator
Doll’s Eyelid
US Patent 20020049023
http://www.medicalartprosthetics.com/content.php?page=galleries&gallery=orbital
Non-Blinking Prosthesis

## Slide 8: Mechanical Wind-Up
User winds mechanism
Blink controller releases one gear at a time
Spring provides quick retraction of lid
Limited blinks per insertion
Open
Closed
Gear
Spring
Shell
1 cm
1.1cm

## Slide 9: Push Solenoid
Eyelid rotates on central axis
Tubular, push type solenoid rotates eyelid
Eyelid “closed” when solenoid powered on
Counterweight opens eyelid
Battery powered with on/off switch
Solenoid
Counterweight
Power Supply
Open
Closed
1 cm
1.1cm

## Slide 10: Pneumatic
Eyelid rotates on central axis
Inflated balloon catheter rotates eyelid
Manually powered with an air bulb
Counterweight opens eyelid
Open
Closed
1.1cm
1 cm
Balloon

## Slide 11: Design Matrix
Criteria | Weight | Pneumatic |  | Wind-Up | Push Solenoid
Feasibility | 1 | 4 | 2 |  | 3
Size | 1 | 3 | 4 |  | 4
Reproducible | 0.3 | 3 | 3 |  | 3
Safety | 0.7 | 4 | 4 |  | 3
Cost | 0.3 | 4 | 3 |  | 2
Risk of Failure | 0.6 | 3 | 1 |  | 2
Appearance | 0.3 | 3 | 3 |  | 3
TOTAL |  | 14.30 | 11.90 |  | 12.45

## Slide 12: Final Design
Open failure mode
Eye glasses conceal air tube
Open
Closed
Closing
300 ms

## Slide 13: Future Work
Finalize Retraction Mechanism
Testing
  Large-scale proof of concept
  Force-weight –retraction balance
Fabrication
  Minimize catheter size
  Minimize deflation time
  Maximize force exerted upon expansion

## Slide 14: References
M.Honda, A.Niimi, M.Ueda“New Orbital Prosthesis With a Blinking Eyelid: Report of a Case” Journal of Oral and Maxillofacial Surgery, Volume 57, Issue 6, Pages 730-733
http://www.sciencedirect.com.ezproxy.library.wisc.edu/science?_ob=MImg&_imagekey=B6WKF-4CRR5T0-F0-1&_cdi=6905&_user=443835&_orig=search&_coverDate=06%2F30%2F1999&_sk=999429993&view=c&wchp=dGLzVlz-zSkzk&md5=3031db4f5557f4f0e54f082b1d6f4444&ie=/sdarticle.pdf
http://www.freepatentsonline.com/20020049023.pdf
http://www.medicalartprosthetics.com/content.php?page=galleries&gallery=orbital

## Slide 15
Questions?