## Slide 1
Team Members: Lacey Halfen, Jessica Hause, Erin Main, Peter Strohm & Fan Wu
Client: Orhan Unal Advisor: Willis Tompkins
Abstract: A recently developed intravenous catheter aims to treat atrial fibrillation under MRI guidance. This device should effectively eliminate the patient’s exposure to X-ray radiation during a traditional intravenous catheterization procedure. The heart phantom design described provides an adequate testing environment for the catheter before animal trials. The dimensions of the phantom’s outer casing fit tightly inside a 13’’ diameter cylindrical coil. The inner contents of the phantom mimic the physiological passage of the catheter into the atria. This internal vasculature was created using flexible and braided tubing, polycarbonate tube inserts and snap grip clamps. Two acrylic hemispheres were combined to resemble the atrial chambers of the heart. An acrylic square with holes of various sizes will allow for high resolution testing in the vicinity of the catheter tip. Initial MRI testing demonstrated that the phantom is completely MRI compatible and is sufficient for testing the catheter under real time MRI imaging.
Final Design
Casing
Pulmonary Veins
Superior Vena Cava
Outlet
Inlets
Drain
Inferior Vena Cava
Catheter entry point with sheath
Heart Chamber
Pulmonary Veins
Inlets
Superior Vena Cava
Outlet
Inferior Vena Cava
Disposable Septum
Catheter puncture from right to left atrium
O-rings
Seal heart chamber
Screws & Bolts
Seal heart chamber
Catheter entry point
Vasculature
Pulmonary Veins
  ½’’ ID tubing
  Braid reinforced tygon
Superior Vena Cava
  1’’ ID tubing
  Tygon
Inferior Vena Cava
  1’’ ID tubing
  Braid reinforced tygon
Figure 5. Braid reinforced tygon tubing used for the pulmonary veins and inferior vena cava.
Problem Statement
Design Constraints
Background Information
Testing
Future Work
This project consists of designing a heart phantom to be used for the initial testing of a solenoid-tipped catheter, which will ultimately be used to treat atrial fibrillation under MRI guidance. The phantom will be used to test the maneuverability of the catheter as well as the high resolution imaging capabilities in the vicinity of the solenoid tip. Vasculature must be transparent and mimic the superior and inferior vena cava and pulmonary veins. The heart chamber must consist of the right and left atrial chambers with a disposable septum between.
Atrial Fibrillation
  Characterized by irregular atrial beating, poor blood flow, increased blood clot risk and disordered electrical signals
  Complications include stroke (blood clots) and heart failure
Pulmonary Vein Isolation
  Catheter Pathway
    Enters femoral vein and travels through inferior vena cava to right atrium
    Puncture through from right atrium to left atrium
  Tissue Ablation
    Irregular impulses originate from pulmonary veins
    Scaring of heart tissue to isolate irregular signals
Client’s Research
  Treatment of atrial fibrillation under MRI guidance using solenoid tipped catheter
Heart Phantom for MRI Guided Interventions
Size
  Diameter 13’’
  Length 13’’
Vasculature
  Pulmonary Veins: inlets
    10 – 20 mm (diameter)
  Superior Vena Cava: outlet
    20 – 30 mm (diameter)
  Inferior Vena Cava: catheter entry point
    20 – 30 mm (diameter)
Heart chamber division
  Thin disposable membrane
MRI compatible (no ferrous materials)
Transparent and leak proof
Leak Testing
  casing, vasculature and heart chamber
Flow Testing
  dynamic pressurized state using pump
  vasculature completely filled with water
MRI Testing
  superimposed images from two coils
    catheter imaging coil
      small field of view
      bright image of catheter
    large imaging coil surrounding phantom
      large field of view
  tracking coil on tip of catheter
  real time tracking with software
    visual feedback during catheterization
Figure 10. MRI image of catheter.
Tubing Connections
Polycarbonate Tubing
Tapered to prevent catheter from catching
Acrylic Wall
Snap Grip Clamps
Tygon Tubing
Fluid Flow & Catheter Pathway
Flow Pathway
Catheter Pathway
Figure 2. Location of catheter during pulmonary vein isolation.
References & Acknowledgements
Replace caulking in heart with softer, better sealing o-rings for quicker removal and replacement
Decrease size of casing to fit into client’s new smaller coil
Flow testing with catheter puncturing septum
Figure 1. Area in left atrium with tissue ablation occurs.
Figure 3. MRI machine that will be used in the testing of the heart phantom.
Figure 6. Heart phantom.
Figure 4. Casing of heart phantom.
Figure 8. Heart chamber.
Figure 9. Fluid flow and catheter pathway through heart chamber.
Figure 7. High resolution block.
High Resolution Block
Visualization of catheter and surrounding vasculature on MRI
  Central hole
    Superior vena cava
  Small holes
    Surrounding vasculature
Mayo Clinic Staff (2007) Heart Disease: Atrial Fibrillation. Retrieved February, 2009, from http://www.mayoclinic.com/
http://www.pic-see.com/MRI_s__CT___Ultrasound.html
http://www.cpmc.org/services/heart/tx/pvai.html
We would like to thank our client Orhan Unal for his help and support during the past two semester, as well as our advisor Willis Tompkins for his guidance.